## Sheet1
Repeatability
0.00999	0.00999	0.00999	0.01000	0.01001	0.01000	0.01001	0.01001	0.01000	0.01001
99.99997	100.00001	99.99998	99.99998	99.99997	99.99998	99.99997	99.99996	99.99999	99.99996
199.99995	200.00000	199.99998	199.99999	199.99996	199.99996	199.99996	199.99997	199.99995	199.99996

Indication error
g	0.01	0.5	1	10	20	50	100	150	200	210	220
mg	0.00	0.00	0.00	-0.01	-0.01	-0.02	-0.01	0.00	0.01	0.03	0.02
mg	0.04	0.05	0.05	0.08	0.10	0.12	0.20	0.31	0.40	0.45	0.47

Certificate 1473 D-K 17296 from 2019-05-10

				Eccentricity			Temp	B press	Rel Hum
				99.99997			19.7	751.7	44.8
				99.99996			20	751.8	47.3
				100.00003
				99.99998
				99.99994

Comments>
This calibration certificate documents the traceability to national standards, which realize the units of measurement according to the International System of Units (SI). The results of this certificate are related to the calibrated object.

For the calibration is used the method of direct comparison with standard weights, in accordance with the document “Guidelines on the Calibration of Non-Automatic Weighing Instruments” EURAMET Calibration Guide No 18 version 4.0 (11/2015). The following tests applies: eccentricity, it determines the difference of indication of the instrument with load in peripheral positions, as opposed to the position in center of the load receptor. Repeatability, to quantify the difference between the results of several weighing ones of the same load when it is deposited several times and of practically identical form on the load receptor and error of indication, considers the performance of the instrument in the total range of measurement

Review in a periodic way the behavior of the balance by means of control with calibrated weights. If the balance is moved to another location after the calibration are likely to alter performance of the balance and may invalidate the calibration. The conformity of the equipment is responsibility of the user according to the use and tolerances established in the processes. Internal adjustment was made to the balance.

Conv Mass correction
210.10051
19.265
0.66325
0.02153
55.46231